## Slide 1: Development of characterization facilities and front-end electronics for SiPM
R. A. Shukla, For SiPM Development Group

## Slide 2: Overview
SiPM Development
SiPM Fabrication
Instrumentation
Characterization
Application

## Slide 3: Silicon Photo-Multiplier (SiPM)
Silicon Photo Multiplier (SiPM) consists of 2-d array of Avalanche Photo Diodes
10 μm – 50 μm
10 μm – 30 μm
Features and Advantages
Operated in Geiger Mode to obtain high Gain ~ 105
Series resistor to quench the Avalanche
High Photon Detection Efficiency ~ 60%
Fast response ~ 100 ps
Immunity to magnetic fields
Compact Size (3mm × 3mm)
Low operating voltage (30-100 V)
Disadvantages:
Limited dynamic range (No. of pixels)
Temperature dependent Gain
Noise – Dark Counts, After-pulses, cross talk etc.

## Slide 4: Micron Resolution Optical Scanner for Silicon Detectors
Indigenously developed tool to characterize optical devices with fine feature size (~10 μm)
To study microscopic uniformity of the response over large area (distinguish between active and dead area)
Very few research laboratories in the world have a facility to carry out such tests with high spatial resolution (1.7 μm)
Active area (Good response)
Dead area (very low response)
Notes: Many of you may be already familiar with SiPM but just for sake of completeness I would like to give a quick Introduction to the device and terminologies used in describing the data.
Large area

## Slide 5: Micron Resolution optical scanner: Why important for SiPM?
For good SiPM , if single pixel output charge = q, then for n pixels, Total output charge = n × q
If pixel to pixel Gain varies, linearity is compromised
Variation of Gain within pixel is also not desirable – induces positional dependence
Number of pixel fired (number of photons)
Charge, Q
All APS’s produce same charge, q1=q2…=q
Variation in pixel to pixel charge

## Slide 6: Beam Focusing and Characterization
LASER beam should be much smaller than 10μm
20x and 50x objectives have been tried
Beam profile was obtained by classic Knife-edge method
Obtaining the sharply focused light beam is the heart of this experiment
10 μm – 50 μm
10 μm – 30 μm
Active area
Dead area

## Slide 7: Characterization of the LASER beam: Knife Edge Method
Collimated beam from LASER (650nm)
50x lens
Knife edge
Detector
( PIN diode)
Beam waist
With 50x lens beam spot size obtained is 1.7 μm at 1σ level
Notes: Because of diffraction, finite beam spot size.
Fwhm = 2.355 sigma

## Slide 8: Fixing target plane position at Focal plane
To scan SiPM : SiPM surface needs to be aligned with Focal plane
A courser position: Imaging capability of online microscope
Exact position: Found out by comparing response of the SiPM at various position along the focal axis
Focal plane

## Slide 9: Test Sample Imaging
Test Sample imaging is a one of the very important feature of the setup
The Experimental setup is as shown in the block diagram
CCD Image of SiPM under test
Imaging capability also allows user to select test area interactively with visual feedback without interfering in the mainstream experiment

## Slide 10: Determination of exact Focal plane
1 p.e events as function of position along focal axis
SiPM is moved along the Beam axis
The plot of the Single pixel triggering events as
function of focal axis position gives the position
of the focal plane.
  1 p.e events starts increasing when SiPM plane
  approaches focal plane

## Slide 11: Transverse Scan of SiPM
Once the focal plane is fixed, SiPM is moved in transverse direction
The whole setup is automated using LabView framework
Typical step size is kept 2 µm and 10000 events are collected at each point at the rate of 1000 events/sec
Typical data size is :
= 125 *125 steps * 10000 events /step
= More than 150 Million triggers
Entire data analysis package is developed with ROOT framework

## Slide 12
RESULTS : 2-D plots ( 1 p.e peak integration)
A typical histogram from active region
A typical histogram from dead region
The contrast in number of events (integration) in 1 p.e events between active and dead region demonstrates excellent sensitivity of our instrument !!

## Slide 13: Scan indicating less sensitive pixel
Usefulness of the was immediately proven when we found a pixel with considerably low gain
One more scan was taken near the bad pixel to confirm the results and for more investigation
More investigation shows that the sensitivity of the Bad pixel is low by ~25%
Publication: “A micron resolution optical scanner for characterization of silicon detectors”,
R. Shukla et. al, Review of Scientific Instruments 85, 023301 (2014); doi: 10.1063/1.4863880

## Slide 14
Instrumentation for SiPM
Amplifier
qADC
TDC
Discriminator
Programmable Power Supply

## Slide 15: Need for PPS with temperature compensation
SiPM’s are being widely used in many high energy physics experiments with large number of channels
Gain of the SiPM varies significantly with temperature, i.e. ~ 5% / °C due to variation in breakdown voltage with ambient temperature
Indoor experiments like CMS at CERN, houses SiPMs in controlled environment thus allowing use of thermo-electric coolers
In outdoor experiments like GRAPES-3, the scintillator detectors are in field
The ambient temperature variation ~ 15°C leading to gain variation of more than 50%
The thermo-electric coolers are not very effective over such large temperature range

## Slide 16
Keep the over-voltage constant
Gain stability ~ 1% required
System Requirements:
  High precision and stable voltage generator
  (~ 6 mV in 0-90 V)
  Accurate temperature sensing (0.1°C)
  Intelligent feedback controller
Temperature compensation concept
Increases by ~ 50 mV/ °C
SiPM Gain α (Applied Voltage – Breakdown Voltage)
Increase by ~ 50 mV/ °C

## Slide 17: Block diagram of the PPS system

## Slide 18: PPS System Calibration and Validation
PPS system was calibrated against standard Keithley pico-ammeter and voltmeter
High precision of ~ 5 mV in output voltage (0-90V) and ~ 0.4 nA in current measurement achieved.

## Slide 19: PPS Testing with SiPM
SiPM was exited with fast laser pulsed and response was recorded with qADC
Complete setup was heated to ~ 35 °C and then allowed natural cooling
The qADC data (charge histogram), temperature, SiPM voltage, SiPM current was recorded continuously
Same exercise was done, with and without temperature compensation

## Slide 20: PPS testing with SiPM: Results
Gain stabilized to ~0.5% over 15°C
Publication:
“Multi-channel programmable power supply with temperature compensation for
silicon sensors”, R. A. Shukla, V. G. Achanta, S. R. Dugad .. et.al
Review of Scientific Instruments, 87, 015114 (2016); doi: 10.1063/1.4940424

## Slide 21: Development of generic modules
Pluggable USB 2.0 Mezzanine Card
A pluggable microcontroller card with full speed USB 2.0 interface has been developed in-house
Sports a versatile I2C interface and general purpose I/Os for user application
Complete software APIs with easy to use I2C library
Used in various application in GRAPES-3 experiment, BARC and CMS experiment
High Speed Amplifier for SiPM (Photo-detectors)
High bandwidth (~ 200 MHz), high gain (~40) amplifier
Very compact form factor; can be integrated into many systems and applications
Compact, high speed amplifier
USB 2.0 Mezzanine card
USB 2.0 Port
I2C Port

## Slide 22: Amplifier + Discriminator – Test data
SiPM Data with direct AFG trigger
SiPM Data with discriminator trigger
Chart: Threshold (mV)
| 12.6 | 24.6 | 42.1 | 60
Rate (Hz) | 100000 | 20000 | 2600 | 150

## Slide 23: Summary
Micron Resolution Scanner has used to characterize SiPMs microscopically
Detailed article with advanced characterization results is underway
Various instrumentation blocks developed in the group will be used for following applications
  Vehicle Monitoring System
  CMS experiment upgrade
  GRAPES-3 experiment
A high speed digitizer development is underway, which will complete the front end instrumentation vertical slice.