## Slide 1: Dipankar Pal, SMIEEE, FIE (I), LMISTE Professor, Dept. of EEE BITS-Pilani, K. K. Birla Goa-Campus Goa-403 726, INDIA
Neural Signal Recording: Opportunities and Challenges
Invited Talk* by
*based on a paper: “Very Low-Noise ENG Amplifier System Using CMOS Technology”,
by R. Rieger, M. Schuettler, D. Pal, C. Clarke, P. Langlois, J. Taylor, and N. Donaldson, published in IEEE Trans. Neural System & Rehabilitation Engg., vol. 14, No. 6, pp. 427-437, December 2006

## Slide 2
Acquisition of bio-signals for advanced medical applications:
ENG-recording to control functional electrical stimulation (FES) prostheses, detection and localization of brain activity.
Acquisition ECG or surface-EMG as part of wearable monitoring system.
These signals are small, on the order of millivolts or less.
Noise and interference are key factors.
Amplification near the recording site is desirable.
Why Record Bio-Signals?

## Slide 3
Advances in CMOS technology
Communication
Low power circuit design
spurred wearable biomedical devices, miniaturized/ highly integrated systems for continuous monitoring.
Crucial building block: sensor interface to pick-up extremely small inputs/ provide preconditioned signal to processing block.
Amplitudes: ~ tens of V to tens of mVs. Frequencies: DC to a few kHz.
Technology-Support

## Slide 4
gain-adjustment during recording
maximum amplification in multi parameter/ multichannel recorders, with matched gain between channels
Choice of input transistor –BJT, MOS, CMOS …….
Noise and input impedance: MOS → very high input impedance,
BJT→ lower noise.
Compromise solution: BJT as lateral structure in CMOS process
Controllable Gain Amplifiers

## Slide 5
In neuroprosthetics research: to use natural ENG to provide sensory feedback to artificial devices.
Neural afferent signals generated by natural sensors within body gives information (like skin contact, force, or limb position).
Can be used in closed-loop neuroprostheses.
Acquisition front ends require further effort like parallel recording-channels: for velocity discrimination.
Interfacing to a live neuron is delicate.
ENG-Recording: A Challenge

## Slide 6
A system for simultaneous recording in 10 channels for velocity classification
Very low-noise (total input referred noise voltage in the bandwidth of 1hz to 5khz ≤ 300nV) and very high gain (~80dB)
A set of 10, ready-to-use by signal processing unit (SPU), dipole outputs
The SPU (not discussed here) consists of digital sections for realizing tripole outputs for EMG cancellation and for velocity classification
Low inter-channel cross-talk
A system for recording ENG from MEC, an extension of conventional nerve cuff, used for selective recording of ENG, classified by action potential velocity.
The design features:
Implantable Multichannel ENG
Recording System

## Slide 7
electrodes
nerve
cuff
Insulating Cuff & Tri-pole
Electrode Assembly

## Slide 8
Impedance Model:
Electrical Circuit
Z0
Ze1
+
-
+
+
-
-
Cuff
IEMG
IENG1
IENG2
Zt2
Zt1
Dipole o/p: Vod1
Ze2
Ze3
Dipole o/p: Vod2
Tripole
o/p: Vot

## Slide 9
ENGs recorded by 11-contact MEC connected to a bank of tripole amplifier system
2-Action Potentials
(of different velocities):

## Slide 10
System Block and SPU Section
(N-2)
2nd
-
rank
amplifiers
1st
-
rank
amplifiers
nerve
electrode
(rings)
insulating
cuff
adder
bandpass
filter
time
delays
output for
one matched
velocity
etc
(N-1)
t
(0)
t
(N-3)
t
t
AC coupling
stage
summers
etc
signal processing unit
(SPU-digital)
: not included

## Slide 11
Basic Preamplifier Architecture
Candidate OTA circuits 
+
¯¯
R1 = 60 k
Filter
R1 =
34 k
C1 =
50 pF

## Slide 12
Device, M | 1,2 | 3 | 4 | 5 | 6 | 7 | 8 | 9-12
W, m | 550 | 51 | 20 | 80 | 160 | 5 | 40 | 10
L, m | 65 | 10 | 4 | 4 | 4 | 10 | 5 | 5
Preamplifier section
Ibias
VDD
VSS
Q1
Q2
M6
M4
M5
M7
M8
M9
M10
M11
M3
M1
M2
M12
Q3
Vin+
Vin-
60 kΩ
34 kΩ
Vout

## Slide 13
V
DD
V
ss
V
in
2/205
2/295
High Valued Active-R for
AC Coupling Stage

## Slide 14
Output noise PSD before / after RC filter (total rms noise in the bandwidth 1Hz-10KHz remains the same)

## Slide 15
Percentage variation of the value of the active resistance from its nominal value (8.2M) as a function of applied voltage compared to a single PMOS transistor of aspect ratio 2/387
-5
-2.5
0
2.5
5
-0.1
-0.05
0
0.05
0.1
Voltage across resistance (V)
Deviation from nominal resistance (%)
Series transistors
Single PMOS transistor

## Slide 16
Device, M | 1-2 | 3-6 | 7 | 8 | 9,10 | 11 | 12 | 13,14 | 15 | 16 | 17,18
W, m | 200 | 5 | 8 | 50 | 6 | 13.5 | 69.3 | 70 | 45 | 15 | 10
L, m | 2 | 35 | 5 | 5 | 45 | 8 | 4 | 5 | 8 | 4 | 5
2nd Rank Amplifier Section
50 kΩ
500 Ω
Ib
M7
M3
M2
Vin+
M1
M5
M4
M9
M8
M6
M10
C1
M13
M12
M11
M15
M16
M17
Vin-
Vout
M14
M18

## Slide 17
ASIC layout. Unidentified structures are additional test circuits

## Slide 18
Die mounted in PGA IC package for testing

## Slide 19
Averaged frequency response of all 10 dipole channels from 5 randomly chosen chips (50 channels in all) compared to nominal simulation

## Slide 20
\
Average input-referred voltage noise taken from 2 channels chosen at random from the same 5 chips used for Fig 14 (10 channels in all)

## Slide 21
Parameter | Specification | Measured
Number of channels (dipoles) | 10 | 10
Power supply | ±2.5 V | ±2.5 V
Power consumption | <50 mW | 24 mW
Circuit area |  | 12 mm2
Midband gain | 10,000 | 10,100
–3 dB frequencies: lower upper | 300 Hz 3.5 KHz | 310 Hz 3.3 KHz
CMRR @ 1KHz | 100 dB | 82 dB
PSRR @ 1KHz VDD VSS | > 40 dB > 40 dB | 42 dB 54 dB
Adjacent channel interference (crosstalk) | > 40 dB | 45 dB
Total input-referred voltage noise density @ 1Hz @ 1kHz | < 20 < 4 | < 20 < 3.8
Total input-referred current noise density @ 1Hz @ 1kHz | < 20 < 2 | < 20 < 1.5
Total input-referred rms voltage noise 1 Hz-5 KHz | < 300 nV | 291 nV
Residual input DC current | < 100 nA | 15.5 nA (+ve. inputs) 20.25 nA (-ve. inputs)
TABLE I: System Specification

## Slide 22
Setup for recording electrically evoked compound action potentials from frog nerve in Ringer’s solution using an 11-contact recording cuff and a hook electrode (for validation purposes).
A polyimide thin-film cuff electrode with eleven contacts, visible on this photograph as dark vertical lines.

## Slide 23
Tripolar recordings of electrically evoked potentials, recorded with 11-contact cuff: stimulation intensity = 0.13 µC. Black bar to the right shows amplitude scale: 50 µV.

## Slide 24
Tripolar recordings of electrically evoked potentials, recorded with 11-contact cuff. The stimulation intensity = 1.01 µC. Black bar to the right shows the amplitude scale: 50 µV.

## Slide 25
Conclusion
Major opportunities in neuroprosthetics-research:
  Use of naturally-occurring ENG to provide sensory feedback to artificial devices.
  Afferent signals from natural sensors can give informations on skin contact, force, and position uable in closed-loop neuroprostheses.
Major challenges in neuroprosthetics-research:
  These applications require stable responses from clinically implanted electrodes. Nerve cuff electrodes report safe implantation for up to 15 years.

## Slide 26
CONCLUSION (Contd.)
A system of implantable MEC with 10-parallel recording channels is presented with features:
  Low noise, highly gain dipole outputs, velocity classification & precise and EMG-cancellation through a subsequent digital SPU unit.
  Challenge of recording extremely small ENG signal in presence of noise is met by a specially designed very low noise preamplifier front-end, that uses lateral pnp in standard CMOS process as optimum choice

## Slide 27
CONCLUSION (Contd.)
A stage to AC couple the preamplifier to subsequent stages is included. Features:
  Removes any preamplifier output offset voltages.
  Defines the lower cut-off frequency of the recording channel (at about 300 Hz).
  Uses a very large active resistance with acceptable linearity.
2nd rank amplifier design with features:
  Cascade of a balanced OTA and a class-AB stage with a capacitor interspersed in-between deciding the upper 3-db cut-off.
Conventional tripoles are known to provide single channel output and also suffer from the inaccuracies involved in EMG cancellation by analogue means.

## Slide 28
“All things must have a usefulness; that is certain. Since electricity must have a usefulness, and we have seen that it cannot be looked for either in theology or in jurisprudence, there is obviously nothing left but medicine”: J. G. Krueger, a Professor in Halle , in 1743
https://archive.org/details/essayonelectrici1785adam: George Adams, Essays on Magnetism, 1785

## Slide 29
REFERENCES
Rieger Robert, “Variable-gain, low-noise amplification for sampling front ends”, IEEE Transaction on Biomedical Circuits and Systems, Vol. 5, No. 3, pp. 253-261, June 2011.
Haugland, M.K., Hoffer, J.A. Slip information obtained from the cutaneous electroneurogram: Application in closed loop control of functional electrical stimulation. IEEE Trans. Rehab. Eng., Vol 2, pp. 29-36, 1994.
Haugland, M.K., Hoffer, J.A., Sinkjaer, T. Skin contact force information in sensory nerve signals recorded by implanted cuff electrodes. IEEE Trans. Rehab. Eng., Vol 2, pp. 18-28, 1994.
B. Popovic and R. B. Stein et al., “Sensory nerve recording for closedloop control to restore motor functions,” IEEE Trans. Biomed. Eng., vol. 40, no. 10, pp. 1024–1031, Oct. 1993.
M. A. Lebedev and M. A. Nicolelis, “Brain–machine interfaces: Past, present and future,” Trends Neurosci, vol. 29, no. 9, pp. 536–546, Sep. 2006.
M. Shojaei-Baghini, R. K. Lal, and D. K. Sharma, “A low-power and compact analog CMOS processing chip for portable ECG recorders,” in Proc. ASSCC, 2005, pp. 473–476.

## Slide 30
R. G. Haahr, S. Duun, E. V. Thomsen, K. Hoppe, and J. Branebjerg, “A wearable “electronic patch” for wireless continuous monitoring of chronically diseased patients,” in Proc. 5th Int. Workshop on Wearable Implantable Body Sensor Networks, 2008, pp. 66–70.
W. Rushton, “A theory of the effects of fibre size in medullated nerves,”J. Physiol, vol. 115, pp. 101–122, 1951.
Haugland, M., Lickel, A., Haase, J., Sinkjaer, T. Control of FES Thumb Force Using Slip Information Obtained from the Cutaneous Electroneurogram in quadriplegic man. IEEE Trans. Rehab. Eng., Vol 7, No. 2, 1999.
Strange, K.D., Hoffer, J.A. Gait phase information provided by sensory nerve activity during walking: Applicability as a state controller feedback for FES. IEEE Trans. Biomed. Eng. Vol 46, No. 7, pp. 797-810, 1999.
Waters, R.L., McNeal, D.R., Faloon, W., Clifford, B. Functional electrical stimulation of the peroneal nerve for hemiplegia: Long term clinical follow-up. Jnl. Bone Joint Surg., 67A, pp. 792-793, 1985.
Continued

## Slide 31
Continued
Haugland, M.K., Sinkjaer, T. Cutaneous whole nerve recordings used for correction of footdrop in hemiplegic man. IEEE Trans. Rehab. Eng., Vol. 3, pp. 307-317, 1995
Schmidt, R.A., Bruschini, H., Tanagho, E.A. Feasibility of inducing micturition through chronic stimulation of sacral roots. Urol., 12, pp. 471-477, 1978.
Rushton, W.A.H, A theory of the effects of fibre size in medullated nerves. J. Physiol. 115:101-122, 1951.
Taylor, J, Donaldson, N and Winter, J. Multiple-electrode nerve cuffs for low velocity and velocity-selective neural recording, Medical and Biological Engineering and Computing, 2004, Vol 42, pp634-643.
M. Steyaert, W. Sansen, C. Zhongyuan, “A Micropower Low-Noise Monolithic Instrumentation Amplifier for Medical Purposes”, IEEE J. Solid-State Circuits, vol. SC-22, no. 6, pp 1163-1168, 1987.
R. Martins, S. Selberherr, F. A. Vaz, “A CMOS IC for portable EEG acquisition systems”, IEEE Trans. Instrumentation and Measurement, vol. 47, no.5, pp. 1191-1196.

## Slide 32
Continued
J. Ji, and K. D. Wise, “An Implantable CMOS Circuit Interface for Multiplexed Microelectrode Recording Arrays”, IEEE J. Solid-State Circuits, vol. 27, no. 3, 1992.
Y Perelman and R Ginosar, 'An integrated system for multichannel neuronal recording with spike/LFP separation and digital output', Proc 2nd International IEEE IEMBS Conference on Neural Engineering, Arlington Virginia, 2006.
R. Rieger, J. Taylor, A. Demosthenous, N. Donaldson, P. Langlois, “Design of a Low-Noise Preamplifier for Nerve Cuff Electrode Recording”, IEEE J. Solid-State Circuits, pp. 1373-1379, vol. 38, no. 8, 2003.
Robert Rieger, Martin Schuettler, Dipankar Pal, Chris Clarke, Peter Langlois, John Taylor and Nick , “Very low-noise ENG amplifier system using CMOS technology”, IEEE Trans. on Neural Systems & Rehabilitation Engineering, Vol. 14, No. 4, pp. 427-437, December 2006.
C. Pflaum, R. R. Riso, G. Wiesspeiner, “Performance of alternative amplifier configurations for tripolar nerve cuff recorded ENG,” Proc. IEEE Int. Conf. Eng. Med. Biol. Soc. (EMBS), vol. 1, pp. 375-376, Amsterdam, Netherlands, 1996.
R. Rieger, J. Taylor, E. Comi et al, Experimental determination of compound action potential direction and propagation velocity from multi-electrode nerve cuffs, Medical Engineering & Physics 26 (2004) 531-534.

## Slide 33
Continued
T. Sinkjaer, M. Haugland, A. Inmann, M. Hansen, K. D. Nielsen, “Biopotentials as command and feedback signals in functional electrical stimulation systems”, Med. Eng. & Phys., vol. 25, pp. 29-40, 2003
E. A. Vittoz, “MOS Transistors Operated in the Lateral Bipolar Mode and Their Application in CMOS Technology”, IEEE J. Solid-State Circuits, vol. SC-18, no. 3, 1983.
J. Erlanger and H. Gasser, ‘The Comparative Physiological Characteristics of Nerve Fibres,’ in Electrical Signs of Nervous Activity. The Eldridge Reeves Johnson Foundation for Medical Physics, Ed. Philadelphia: University of Pennsylvania Press, pp. 34-78, 1937.
R Geiger, P Allen and N Strader, 'VLSI design techniques for analog and digital circuits', McGraw Hill, 1990, chapter 5.
P Allen & D Holberg, 'CMOS analogue circuit design', Oxford University Press, 2002, chapter 5.
T. Stieglitz, H. Beutel, M. Schuettler, and J. U. Meyer, ‘Micromachined Polyimide-Based Devices for Flexible Neural Interfaces’, Biomedical Microdevices, vol. 2, no. 4, pp. 283-294, 2000.

## Slide 34
Thank you