## Slide 1
Production and measurement
Techniques of Very low temperature
Anil Kumar
Department of Condensed Matter Physics and Materials Science,
Tata Institute of Fundamental Research, Mumbai-400005, India

## Slide 2
Plan
Introduction
Adiabatic Demagnetization Refrigeration (ADR)
Dilution Refrigerator System (DRS)
Thermometry
Summary
Major Results and Future Plans

## Slide 3
What do we mean by Very Low Temperature (VLT)?
< 1K is called VLT
T(K) = T(ºC) + 273.15

## Slide 4
~ 300 K Room Temperature
2.7 K Temperature of Universe (CMBR)
3.2 K Boiling point of Liquid Helium 3 (3He)
4.2 K Boiling point of Liquid Helium 4 (4He)
77.4 K Boiling point of Liquid Nitrogen
~184 K Coldest temperature on earth (-89.2 C,Vostok, Antarctica)
273 K Zero degree Celsius (Ice point)
0.00001 K Nuclear cooling (Copper lattice + electrons)
0.005 K Commercial Refrigerator based on
Dilution 3He in 3He-4He mixture
0.3 K Pumped 3He bath
1.2 K Pumped 4He bath
0 K Absolute Zero (cannot be attained)
≈

## Slide 5: Historical Development of Refrigeration
1840
1860
1880
1900
1920
1940
1960
1980
2000
10-4
10-3
10-2
10-1
100
101
102
103
O2
N2
H2
4He
3He
Low T
Ultra-low T
10-5
Magnetic refrigeration,
electronic mag. moments
Magnetic refrigeration,
nuclear mag. moments
3He-4He Dilution fridge
Faraday,
chlorine 1823
Below 10 mK the dominant cooling principle is the demagnetization of a nuclear spins of a metal.
Down to 10 mK. The cooling of a gas by letting it do work against a force during an expansion. The gas employed may be a conventional gas or 3He atoms mixed in liquid 4He.
Kamerlingh-
Onnes
10-6
TEMPERATURE (K)
YEAR
1908

## Slide 6
In your chai, the most energetic chai molecules escape from the cup and come off as steam. When they do this, they take away more than their share of heat, and the atoms left behind in the cup are colder because they have lost energy.
For Low Temp, working substance is Helium gas.
BEGIN WITH EVOPORATIVE COOLING
Liquid Helium Cryostat

## Slide 7: Evaporative cooling and its limitations

## Slide 8: Refrigeration Below 1.5K
Adiabatic Demagnetization Refrigerator(ADR)
(4.2K to 70mK)
Dilution Refrigerator System(DRS)
(1.5K to 5mK)

## Slide 9: History Adiabatic Demagnetization Refrigerator(ADR)
Cooling to milli-Kelvin temperatures by effecting electron spins was proposed by.
– P. Debye 1926. – W.F. Giauque 1926
The first practical demonstration was by. – Giauque, MacDougall and De Haas, Wiersma, Kramers, 1933. – Kurti and Simon 1934

## Slide 10: Adiabatic Demagnetization

## Slide 11: Entropy Change of Salt pill as a Function of Temperature

## Slide 12
Paramagnetic materials

## Slide 13
ADR Inserts Schematic Diagram

## Slide 14: Adiabatic Demagnetization Refrigerator (ADR)
Precool ADR Insert in Cryostat with liquid N2 to 77.2K
Cool to 4.2K with Liquid He
Pumped 1K Pot to cool LTS stage to 1.5K
Isothermal Mag and Adiabatic Demag Of Paramagnetic Salt Pill to reach 70 miliKelvin.
ADR Insert
Cryostat
Measurements Electronics
Magnet Insert
(LTS) Low Temperature Stages Of ADR

## Slide 15: Low Temperature Stages Of ADR
AC Resistivity
Sample Holder
Sensors and Heater
Salt Pill
AC Susceptibility Coil
1K Pot
1K POT
Heat Switch
Kevlar thread
Thermal Isolation
1K Plate
Sorb with Heater

## Slide 16
Standard Samples Measured in ADR
I=10µA

## Slide 17: History Dilution Refrigerator System(DRS)
Cooling to milli-Kelvin temperatures by dilution refrigerator was proposed by.
– H. London 1951. – H. London, G.R. Clarke, and E. Mendoza 1962
The first practical demonstration was by. – Das, DeBruyn, & Taconis at Leiden University 1965

## Slide 18
Cooling to Low Temperature
Lq 4He B.P 4.2 K pumping bath ~ 1.2 K (Vapour pressure 1 mbar)
Lq 3He B.P 3.2 K pumping bath ~0.3 K (Vapour pressure ~0.01 mbar)
Dilution Refrigerator (3He/4He mixture) ~ 0.005 K (Dilution cooling)
0
0.50
0.75
1.00
0.25
1.0
2.0
Superfluid
3He/4He
Normal Mixture 3He/4He
Phase separation line
l line
Two phase Region
3He Concentration
T(K)
Con. 3He floats on the Dil. Phase 3He/4He
0.867K

## Slide 19
Cooling done by moving 3He atoms from liquid to gas-like phase. To remove 3He atoms from mixing chamber one needs a still around 0.6 to 0.8K. From still 90% of the outcoming gas is 3He since vapour pressure of 4He is only 5% of that of 3He.
P a ( TS XS – TMC XMC)~2.2 KPa
(1 KPa~1.25 m liq. 3He) acts like a pump to pull 3He atoms.
Heat Exchangers are inserted between the inflowing and outflowing liquid 3He streams to improve the efficiency of the cooling cycle of DRS.
+ 93.5% 4He

## Slide 21: Micro-Kelvin Set-up
DRS Insert
Gas handling system
Pure He4 Gas
Pure Helium 3
Controller Panel
Copper
Nuclear Stage

## Slide 22: Instrument Housing

## Slide 23: Dilution Refrigerator (TIFR) Specifications

## Slide 24
What is a Thermometer?
Thermometers measure temperature, by using material property that changes when heated or cooled. In a mercury or alcohol thermometer the liquid expands as it is heated and contracts when it is cooled, so the length of the liquid in the column is longer or shorter depending on the temperature.

## Slide 25
Thermometry at VLT in our Lab
Low Temperature Resistance sensors used in our systems:-
  RuO2 sensors ( 50mK to 4.2K, H=9T)
  Carbon composite sensors SPEER (10mK to 4K, H=5T)
  Response time is very large at low temperatures(<50mK)~mins.
2. Magnetic susceptibility sensors:
  Cerium Magnesium Nitrate (Paramagnetic Salt) 5mK – 4.2K
SQUID based Noise Thermometer:

## Slide 26: Thermometry 1
RuO2 Temperature Sensor
SPEER carbon Temperature Sensor
Temperature monitored at 1 K stage, Still using RuO2 resistance sensor
(saturates below 50 mK)
Temperature monitored at Mixing Chamber using SPEER Carbon resistance
sensor (saturates below 10 mK)

## Slide 27
Down to 4 mK at the mixing chamber using a CMN (Cerium Magnesium Nitrate) mutual inductance thermometer – initially calibrated against a SPEER resistance sensor
CMN requires a calibration every run at 4.2 K
Thermometry 2

## Slide 28
Noise thermometry is based on the fundamental Nyquist formula,
Where U2 is the voltage fluctuations caused by the thermal agitation of electric charges in an unbiased conductor to thermodynamic temperature T.
4.2 K to 1 mK
Tref= Reference Temperature of Liquid Helium 4.2K
S0(Tref)= Noise Spectrum @ 4.2K
S0 (T)= Noise spectrum which we are measuring.

## Slide 30
Summary
4.2K
Liquid Helium
1.5K
Pumping Over
Liquid Helium
~ 0.1K (100mK)
Adiabatic Demagnetization Refrigerator
(ADR)
~ 0.005K (5mK)
Dilution Refrigerator System
(DRS)
Various thermometers suitable for VLT

## Slide 31
Major Result and Future Plans
Developed DC magnetometer in Dilution Refrigerator System for use in Micro-Kelvin Temperature Range . Discovery of SC in Bi.
Development of DC magnetometer in Adiabatic Demagnetization Refrigerator (ADR) is in progress
Development of SQUID sensor based resistance measurement in Micro-Kelvin Temperature range is also in progress

## Slide 32: Acknowledgement
Prof. S Ramakrishnan, Mr. Sannabhadti , Students, Scientific staff and Central workshop

## Slide 33
Thank you