## Slide 1: Closed Loop Gas System
R.R. Shinde
Behalf of Gas system team
DHEP

## Slide 2: Outline
Introduction
Gas system
Components of Closed loop gas system
Operations in the Closed loop gas system
Upgradation
Problem faced
Present status
Future plan

## Slide 3: Introduction
The RPC, Active element in the INO experiment is a gaseous detector.
Mode of operation is determined by mixture of gas flowing through the RPCs.
Avalanche mode: R134a+Isobutane+SF6
Streamer mode: R134a+Argon+Isobutane
Gas system is required to circulate proper gas mixture in the Detector

## Slide 4: System at Glance
Parameters | mini- ICAL | E-ICAL
Size | 4 mtr x 4 mtr | 8 mtr x 8 mtr
No of Layers | 10 | 20
No of RPCs Per Layer | 2 | 16
Total No. of RPCs | 20 | 320
Size of gas gap in the RPC | 184.5*174*0.2 cm3 | 184.5*174*0.2 cm3
Gas Volume of Each RPC | 6.42 Ltrs | 6.42Ltrs
Total Gas Volume | 128.4 Ltrs | 2054.4 Ltrs
Mode of RPC Operation | Avalanche | 
Required gases | R134a (95.2%) | 
 | Iso Butane (4.5%) | 
 | SF6(0.3%) | 

## Slide 5: Gas system
There are 3 types of gas systems
Open ended gas system: No reuse of the output gas mixture from the detector
Open loop gas system: collect the output gas mixture from the detector and segregate them and reuse in the system
Closed loop gas system: Collect the output gas mixture from the detector, purify it and recirculate it.
We have an open ended system and a Closed loop gas system operational in Mumbai and Madurai.

## Slide 6: Important components of the CLS
Gases
Mass Flow Controllers
Storage Tanks
Pressure sensors
Purification System with Radical remover and molecular sieves
Input and output header
Capillaries as flow resistors
PLC based control and monitoring system
Diaphragm Pump
Vacuum pump
Valves
Compressor
Hygrometer

## Slide 7: Operations in the Closed loop gas system
Closed Loop
Open End
Refilling
Vacuum
Regeneration of molecular Sieves
Manual ON-OFF

## Slide 8: Line Diagram of CLS
RGA
RGA

## Slide 9: Upgradation
Automation of all operation has been carried out
Interlocks are added in the system
Alert system (Siren) added

## Slide 10: Problem faced
Presence of air detected
More number of top up cycles
Rise in current in the RPCs

## Slide 11: Diagnosis Tools
Snoop Solution
Gas Leak Detector (SP220)
Gas Chromatograph
RGA SRS300
Pressure Profiles

## Slide 12: Diagnosis Tools
Chart: RGA_CLS_Oulet_Header_ 1.0e-6
GC Spectrum

## Slide 13: CLS Parameters performance

## Slide 14: Current Status
Intensive leak arrest task completed.
Operational fine tunning is going on
Initial results shows Gas Mixture top-up happens after 3.5 days.
Evaluation of leak arrest based on RPC performance is in progress

## Slide 15: Future plan
Replacement of faulty pneumatically operated valves
Implementation of Capillary of small length for all RPCs
Improve the operation and control philosophy of E-ICAL based on experience at mini-ICAL

## Slide 16: Working forces of Present gas system team
Prof. Gobinda Majumder
Dr. Manish Shah (BARC)
Shri. Piyush Verma
Shri. Santosh S. Chavan
Shri. Vishal V. Asgolkar
Shri. Ganesh K. Ghodke
Umesh Shas
Kirti Prakash Sharma
Ravindra R. Shinde
Shri. K.C. Ravindran
&
Dr. Satyanarayana Bheesette

## Slide 17: Thank You